
Introductory theory

[image: image1.png]



Introduction

Few methods of chemical analysis are truly specific to a particular analyte. It is often found that the analyte of interest must be separated from the myriad of individual compounds that may be present in a sample. As well as providing the analytical scientist with methods of separation, chromatographic techniques can also provide methods of analysis.

Chromatography involves a sample (or sample extract) being dissolved in a mobile phase (which may be a gas, a liquid or a supercritical fluid). The mobile phase is then forced through an immobile, immiscible stationary phase. The phases are chosen such that components of the sample have differing solubilities in each phase. A component which is quite soluble in the stationary phase will take longer to travel through it than a component which is not very soluble in the stationary phase but very soluble in the mobile phase. As a result of these differences in mobilities, sample components will become separated from each other as they travel through the stationary phase.

Techniques such as H.P.L.C. (High Performance Liquid Chromatography) and G.C. (Gas Chromatography) use columns - narrow tubes packed with stationary phase, through which the mobile phase is forced. The sample is transported through the column by continuous addition of mobile phase. This process is called elution. The average rate at which an analyte moves through the column is determined by the time it spends in the mobile phase. 



Distribution of analytes between phases

The distribution of analytes between phases can often be described quite simply. An analyte is in equilibrium between the two phases;

Amobile[image: image2.png]


 Astationary
The equilibrium constant, K, is termed the partition coefficient; defined as the molar concentration of analyte in the stationary phase divided by the molar concentration of the analyte in the mobile phase.

The time between sample injection and an analyte peak reaching a detector at the end of the column is termed the retention time (tR ). Each analyte in a sample will have a different retention time. The time taken for the mobile phase to pass through the column is called tM.

[image: image3.png]i)

|eubis 1018190

Time




A term called the retention factor, k', is often used to describe the migration rate of an analyte on a column. You may also find it called the capacity factor. The retention factor for analyte A is defined as;

k'A = t R - tM / tM
t R and tM are easily obtained from a chromatogram. When an analytes retention factor is less than one, elution is so fast that accurate determination of the retention time is very difficult. High retention factors (greater than 20) mean that elution takes a very long time. Ideally, the retention factor for an analyte is between one and five.

We define a quantity called the selectivity factor,  , which describes the separation of two species (A and B) on the column;

 = k 'B / k 'A
When calculating the selectivity factor, species A elutes faster than species B. The selectivity factor is always greater than one. 


Band broadening and column efficiency

To obtain optimal separations, sharp, symmetrical chromatographic peaks must be obtained. This means that band broadening must be limited. It is also beneficial to measure the efficiency of the column.

The Theoretical Plate Model of Chromatography
The plate model supposes that the chromatographic column is contains a large number of separate layers, called theoretical plates. Separate equilibrations of the sample between the stationary and mobile phase occur in these "plates". The analyte moves down the column by transfer of equilibrated mobile phase from one plate to the next.

[image: image4.png]-The column

Theoretical
Dlate





It is important to remember that the plates do not really exist; they are a figment of the imagination that helps us understand the processes at work in the column.They also serve as a way of measuring column efficiency, either by stating the number of theoretical plates in a column, N (the more plates the better), or by stating the plate height; the Height Equivalent to a Theoretical Plate (the smaller the better).

If the length of the column is L, then the HETP is

HETP = L / N
The number of theoretical plates that a real column possesses can be found by examining a chromatographic peak after elution;

[image: image5.png]



where w1/2 is the peak width at half-height.

As can be seen from this equation, columns behave as if they have different numbers of plates for different solutes in a mixture.

The Rate Theory of Chromatography
A more realistic description of the processes at work inside a column takes account of the time taken for the solute to equilibrate between the stationary and mobile phase (unlike the plate model, which assumes that equilibration is infinitely fast). The resulting band shape of a chromatographic peak is therefore affected by the rate of elution. It is also affected by the different paths available to solute molecules as they travel between particles of stationary phase. If we consider the various mechanisms which contribute to band broadening, we arrive at the Van Deemter equation for plate height;

HETP = A + B / u + C u 
where u is the average velocity of the mobile phase. A, B, and C are factors which contribute to band broadening.

A - Eddy diffusion
The mobile phase moves through the column which is packed with stationary phase. Solute molecules will take different paths through the stationary phase at random. This will cause broadening of the solute band, because different paths are of different lengths.

B - Longitudinal diffusion
The concentration of analyte is less at the edges of the band than at the center. Analyte diffuses out from the center to the edges. This causes band broadening. If the velocity of the mobile phase is high then the analyte spends less time on the column, which decreases the effects of longitudinal diffusion.

C - Resistance to mass transfer
The analyte takes a certain amount of time to equilibrate between the stationary and mobile phase. If the velocity of the mobile phase is high, and the analyte has a strong affinity for the stationary phase, then the analyte in the mobile phase will move ahead of the analyte in the stationary phase. The band of analyte is broadened. The higher the velocity of mobile phase, the worse the broadening becomes.

Van Deemter plots
A plot of plate height vs. average linear velocity of mobile phase.

[image: image6.png]Plate height

Atypical Van Deemter plot

Optirnur velocity

Minirum plate

height

Mobile phase velocity




Such plots are of considerable use in determining the optimum mobile phase flow rate.



Resolution

Although the selectivity factor, , describes the separation of band centres, it does not take into account peak widths. Another measure of how well species have been separated is provided by measurement of the resolution. The resolution of two species, A and B, is defined as

[image: image7.png]Ro 2[(tg)g - (tg),]
Wyt Wy




Baseline resolution is achieved when R = 1.5

It is useful to relate the resolution to the number of plates in the column, the selectivity factor and the retention factors of the two solutes; 

[image: image8.png]



To obtain high resolution, the three terms must be maximised. An increase in N, the number of theoretical plates, by lengthening the column leads to an increase in retention time and increased band broadening - which may not be desirable. Instead, to increase the number of plates, the height equivalent to a theoretical plate can be reduced by reducing the size of the stationary phase particles.

It is often found that by controlling the capacity factor, k', separations can be greatly improved. This can be achieved by changing the temperature (in Gas Chromatography) or the composition of the mobile phase (in Liquid Chromatography).

The selectivity factor, , can also be manipulated to improve separations. When  is close to unity, optimising k' and increasing N is not sufficient to give good separation in a reasonable time. In these cases, k' is optimised first, and then  is increased by one of the following procedures: 

1. Changing mobile phase composition 

2. Changing column temperature 

3. Changing composition of stationary phase 

4. Using special chemical effects (such as incorporating a species which complexes with one of the solutes into the stationary phase)

[image: image9.png]GasjChiomatography;



[image: image10.png]



Introduction

Gas chromatography - specifically gas-liquid chromatography - involves a sample being vapourised and injected onto the head of the chromatographic column. The sample is transported through the column by the flow of inert, gaseous mobile phase. The column itself contains a liquid stationary phase which is adsorbed onto the surface of an inert solid.

Have a look at this schematic diagram of a gas chromatograph:

[image: image11.png]Injector

Flow
contraler P2
i et
'Column] Detector
Column oven

Carrier gas







Instrumental components

Carrier gas 

The carrier gas must be chemically inert. Commonly used gases include nitrogen, helium, argon, and carbon dioxide. The choice of carrier gas is often dependant upon the type of detector which is used. The carrier gas system also contains a molecular sieve to remove water and other impurities. 

Sample injection port 

For optimum column efficiency, the sample should not be too large, and should be introduced onto the column as a "plug" of vapour - slow injection of large samples causes band broadening and loss of resolution. The most common injection method is where a microsyringe is used to inject sample through a rubber septum into a flash vapouriser port at the head of the column. The temperature of the sample port is usually about 50C higher than the boiling point of the least volatile component of the sample. For packed columns, sample size ranges from tenths of a microliter up to 20 microliters. Capillary columns, on the other hand, need much less sample, typically around 10-3 L. For capillary GC, split/splitless injection is used. Have a look at this diagram of a split/splitless injector;

[image: image12.png]The split / splitless injector

7

=+ Septum purge outiet

Carrier gas
inlet

L+ Split outlet

Heated metal block —

Vapourisation charmber
Glass liner-

Calumn





The injector can be used in one of two modes; split or splitless. The injector contains a heated chamber containing a glass liner into which the sample is injected through the septum. The carrier gas enters the chamber and can leave by three routes (when the injector is in split mode). The sample vapourises to form a mixture of carrier gas, vapourised solvent and vapourised solutes. A proportion of this mixture passes onto the column, but most exits through the split outlet. The septum purge outlet prevents septum bleed components from entering the column.

Columns
There are two general types of column, packed and capillary (also known as open tubular). Packed columns contain a finely divided, inert, solid support material (commonly based on diatomaceous earth) coated with liquid stationary phase. Most packed columns are 1.5 - 10m in length and have an internal diameter of 2 - 4mm.

Capillary columns have an internal diameter of a few tenths of a millimeter. They can be one of two types; wall-coated open tubular (WCOT) or support-coated open tubular (SCOT). Wall-coated columns consist of a capillary tube whose walls are coated with liquid stationary phase. In support-coated columns, the inner wall of the capillary is lined with a thin layer of support material such as diatomaceous earth, onto which the stationary phase has been adsorbed. SCOT columns are generally less efficient than WCOT columns. Both types of capillary column are more efficient than packed columns.

In 1979, a new type of WCOT column was devised - the Fused Silica Open Tubular (FSOT) column;

[image: image13.png]Cross section of a Fused Silica Open Tubular Column

Polyiride coating
Fused silica tube
Chernically bonded stationary phase




These have much thinner walls than the glass capillary columns, and are given strength by the polyimide coating. These columns are flexible and can be wound into coils. They have the advantages of physical strength, flexibility and low reactivity.

Column temperature
For precise work, column temperature must be controlled to within tenths of a degree. The optimum column temperature is dependant upon the boiling point of the sample. As a rule of thumb, a temperature slightly above the average boiling point of the sample results in an elution time of 2 - 30 minutes. Minimal temperatures give good resolution, but increase elution times. If a sample has a wide boiling range, then temperature programming can be useful. The column temperature is increased (either continuously or in steps) as separation proceeds.

Detectors
There are many detectors which can be used in gas chromatography. Different detectors will give different types of selectivity. A non-selective detector responds to all compounds except the carrier gas, a selective detector responds to a range of compounds with a common physical or chemical property and a specific detector responds to a single chemical compound. Detectors can also be grouped into concentration dependant detectors and mass flow dependant detectors. The signal from a concentration dependant detector is related to the concentration of solute in the detector, and does not usually destroy the sample Dilution of with make-up gas will lower the detectors response. Mass flow dependant detectors usually destroy the sample, and the signal is related to the rate at which solute molecules enter the detector. The response of a mass flow dependant detector is unaffected by make-up gas. Have a look at this tabular summary of common GC detectors:

	Detector
	Type
	Support gases
	Selectivity
	Detectability
	Dynamic range

	Flame ionization (FID)
	Mass flow
	Hydrogen and air
	Most organic cpds.
	100 pg
	107

	Thermal conductivity (TCD)
	Concentration
	Reference
	Universal
	1 ng
	107

	Electron capture (ECD)
	Concentration
	Make-up
	Halides, nitrates, nitriles, peroxides, anhydrides, organometallics
	50 fg
	105

	Nitrogen-phosphorus
	Mass flow
	Hydrogen and air
	Nitrogen, phosphorus
	10 pg
	106

	Flame photometric (FPD)
	Mass flow
	Hydrogen and air possibly oxygen
	Sulphur, phosphorus, tin, boron, arsenic, germanium, selenium, chromium
	100 pg
	103

	Photo-ionization (PID)
	Concentration
	Make-up
	Aliphatics, aromatics, ketones, esters, aldehydes, amines, heterocyclics, organosulphurs, some organometallics
	2 pg
	107

	Hall electrolytic conductivity
	Mass flow
	Hydrogen, oxygen
	Halide, nitrogen, nitrosamine, sulphur
	 
	 


[image: image14.png]The Flame lonisation Detector

]
Collector electrode
Flame ignition
col 4300V
Polarising vohage
P | S

Colurnn




The effluent from the column is mixed with hydrogen and air, and ignited. Organic compounds burning in the flame produce ions and electrons which can conduct electricity through the flame. A large electrical potential is applied at the burner tip, and a collector electrode is located above the flame. The current resulting from the pyrolysis of any organic compounds is measured. FIDs are mass sensitive rather than concentration sensitive; this gives the advantage that changes in mobile phase flow rate do not affect the detector's response. The FID is a useful general detector for the analysis of organic compounds; it has high sensitivity, a large linear response range, and low noise. It is also robust and easy to use, but unfortunately, it destroys the sample.

