TOPIC 3:

Equipment Acquisition and Risk Management

CONTENTS:

1. Equipment Acquisition

1.1 Process of Acquiring Technology

1.1.1 Overall Process

1.1.2 Pre-purchase Evaluation

1.2 Acquisition Process Strategies

1.2.1 Cost of Ownership

1.2.2 Utilization Information

1.2.3 Life-Cycle Cost Analysis

1.2.4 Equipment Evaluation

1.2.5 The Conditions of Sale Document

1.2.6 Making the Award

1.2.7 Acceptance Testing

1.3 Clinical Team Requirements

2. Risk Management

2.1 Introduction

2.2 Reducing Risk

2.3 Quality Improvement

2.4 Codes and Standards

1.
EQUIPMENT ACQUISITION

Objective: Improved performance at reduced cost

1.1
Process of Acquiring Technology

1.1.1
Overall Process

Typically, medical device systems will emerge from strategic technology planning and technology assessment processes as “required and budgeted” needs.  The following process should be followed:

1. A need analysis should be conducted, reaffirming clinical needs and device-intended applications. 

2. Material management and clinical engineering personnel should ensure this item is a candidate for centralized or coordinated acquisition of similar equipment with other hospital departments.

3. Written pre-purchase evaluation guidelines should be followed.

4. Purchasing takes place, wherein final equipment negotiations are conducted and purchase documents are prepared and a purchase order is made.

1.1.2`Pre-purchase Evaluation

Typical hospital pre-purchase evaluation include the following steps:

1. Analysis of needs and development of a specification list.

2. Developing a vendor list and requesting proposals.

3. Analysing proposals, site planning.

4. Evaluating samples.

5. Selecting finalists.

6. Make the award.

7. Delivery and installation.

8. Acceptance testing

Formal request for proposals (RFPs) from potential equipment vendor are required for intended acquisitions whose initial or life-cycle cost exceeds a certain threshold, e.g., RM100,000.00

1.2
Acquisition Process Strategies

1.2.1
Cost of Ownership

The cost of ownership encompasses all direct and indirect expenses associated with medical equipment over its lifetime.  

It expresses the cost factors of medical equipment for both the initial price of the equipment (which typically includes the equipment, its installation, and initial training cost) and support over the long-term. Long-term costs include ongoing training, equipment service, supplies, connectivity, upgrades, and other costs. 

The full range of cost of ownership factor includes: 

· technology assessment and acquisition costs, 

· operation and maintenance costs (installation, supplies, downtime, training, spare parts, test equipment and tools , and depreciation)

It is estimated that the purchase price represents only 20% of the life-cycle cost ownership.

1.2.2
Utilization Information

When conducting needs analysis (Step 1 as in pre-purchase evaluation), actual utilization information from the organization’s existing similar devices can be very helpful.

One method of measuring equipment utilization is to match product to its intended function and to measure daily (if necessary) the equipment ‘s actual utilization. Equipment functional requirements to meet the demands of clinical procedures are taken into account.  

The following is a simplified description of how this utilization information model is developed. For imaging procedures: knowing devices needed and minutes needed (annually) to accomplish certain procedures, and based on the hospital’s past case mix, the staff can determine equipment needs matched to new/replacement functional capacities. Analysis includes the following information:

· From the equipment inventory, equipment installation dates and condition codes [1(new) to 5(scrap), as determined by technical service appraisal and service expenses]

· From the procedure mix, like procedures are grouped by modality.

· Procedure/equipment  volumes information is retrieved from the central billing system.

· Operational constraints are determined (the total period of time equipment is available during the day). For example , comparing a Radiographic and Fluoroscopic (R and F) x-ray room’s ideal utilization versus a straight radiographic  (Rad) room’s ideal utilization is identified in table 1.

Table 1: R and F Room vs. Rad Utilization

	Exams
	(hrs/day)
	Days per year
	Annual minutes available

	Fluoro – am only
	4 x (60 min/hr)
	260
	62,400

	Rad – pm only
	6 x (60 min/hr)
	260
	93,600

	Rad room – full day
	10 x (60 min/hr)
	260
	156,000


Functional needs are than matched to certain equipment (e.g., for myelogram, a certain x-ray table and other specific accessories may be required.

1.2.3
Life-Cycle Cost Analysis

Life-cycle cost analysis is a tool used during technology planning, assessment or acquisition either to compare high-cost, alternative means for providing a service, or to determine whether a single project or technology has a positive or negative economic value.

The strength of life cycle-cost analysis is that it examines the cash flow impact of an alternative over its entire life, instead of focusing solely on initial capital investments.

Life cycle cost analysis facilitates comparisons between projects or technologies with large initial cash outlays and those level outlays and inflows overtime.

It is most applicable to complex, high-cost choices among alternative technologies, new services, and different means for providing a given service.

It is useful for decisions that are too complex and ambiguous for experience and subjective judgment alone. It helps decision makers perceive and include costs that often are hidden or ignored, and that may otherwise invalidate results.

The most powerful life cycle cost technique is Net Present Value (NPV) analysis, which explicitly accounts for inflation and foregone investment opportunities by future cash flows in present dollars.

A hospital might use LCC analysis as a key factor in such important decision  as:

· To replace/rebuild or buy/lease medical imaging equipment;

· To purchase equipment with a greater initial capital outlay and lower operating costs or to choose an alternative with a lower initial outlay but higher operating costs;

· To acquire and install equipment or contract with a mobile equipment vendor to provide the service  when beginning a major new service (e.g., MRI service);

· To acquire a costly, specialized clinical laboratory instrument or continue to send tests to a reference lab;

· To stay with hard-copy storage of radiology films or move to electronic image archiving;

· To continue to lease a complete telephone system or buy a new one.

Kinds of cost captured in life cycle cost analysis include the following:

· decision making costs,

· planning agency / Certificate of Needs costs (if applicable)

· financing,

· initial capital investment cost

· facility changes,

· life cycle maintenance,

· repair costs,

· personnel costs, and others.

Summary of Steps for LCC/NPV analysis

1. Map out all the factors related to each alternative to gain a complete understanding of each alternative’s impact on the status quo.  This first step is perhaps the most crucial.

2. Choose a time frame, or lifecycle, of the decision.

3. Determine the actual incremental cash flows associated with the quantifiable elements of each alternative.

4. Choose a discount factor.

5. Determine the net present value (or cost) for each alternative.

6. Compare the alternatives and use the comparison as input for final choice.

7. Perform sensitivity analysis using alternative assumptions.

8. Assess the likely impact of all identified subjective factors.

1.2.4
Equipment Evaluation

The process provides information that can be used to screen unacceptable performance by either the vendor or the equipment before it becomes a hospital problem.

The pre-purchase evaluation process consists of technical, clinical, financial, and operational aspects.  These aspects are evaluated during technology assessment process.

Clinical Engineer/Technology Manager should be familiar with emerging and evolving technologies and can translate the clinical needs of users into an effective and comprehensive bids specification document.  

The specification document should be clear to facilitate a competitive bidding environment and a valid comparison of vendors and their product and sets the whole equipment evaluation and selection process into motion.

Validation criteria for key elements should be addressed. The key elements to be addressed are:

· System configuration,

· Extent of facility preparation and operational disturbance,

· Equipment performance requirements,

· User and service provider training,

· Warranty

· Documentation,

· Delivery schedule, and 

· Implementation plan.

Cost of service support and price for future upgrades should be locked in.

The returns for RFP’s from the vendor, the CE department should be responsible for evaluating technical responses, while the materials management/finance department should evaluate the financial responses.

In translating clinical needs into specification list, key features or must-have attributes of desired device are identified.  In practice, hospital should develop a must have list and a list of extras.  The extras list includes features that may tip the decision in favor of one vendor, all other factors being even.

The extras list includes features that may tip the decision in favor of one vendor, all other factors being even.  These specification lists are sent to the vendor and are effective in a self-elimination process that results in saving time for the hospital.

Once the must have attributes have been satisfied, the remaining candidate devices are evaluated technically and extra are considered.  This is accomplished by assigning a weighting factor (e.g., 0 to 5) to denote the relative importance of each desired attribute.

The relative ability of each device to meet the defined requirements are then rated.  Consider the following examples of attributes:

· Accuracy and repeatability.

· Ease-of-use.

· Reliability.

· Expected user’s skill level.

· Serviceability and warranty.

· Performance.

· Compatibility and interchangeability.

· Upgradability or expandability

· Safety

· Cost-effectiveness.

The RFP therefore, should provide detail description of the must have and the extras, the cost of service and upgrades, and should identify a recourse for vendor deficiencies.

1.2.5
The Conditions of Sale Document

One strategy that strengthens the acquisition process is the Condition of Sale document.  This multifaceted document integrates equipment specifications, performance, installation requirements and follow-up services.

The Conditions of Sale ensures that negotiations are completed before a purchase order is delivered and each participant is in agreement about the product to be delivered.

It also specifies the codes and standards having jurisdiction over that equipment.  May include provision for future modification of equipment, compliance with standards under development and provision for software upgrades.

In summary, all powered equipment purchases should be subjected to the terms and conditions as outlined in the Conditions of Sale.  

1.2.6
Making the Award

Standard purchase orders that include the Conditions of Sale for medical equipment are usually used to initiate order.  In addition to current facility conditions are addressed as a minimum:

· Installation and approval requirements, responsibilities, and timetable.

· Payment, assignment, and cancellation.

· Software requirements and updates (e.g., software availability, licensing fees).

· Documentation.

· Clinical and technical training.

· Acceptance testing (facility and vendor).

· Warranty, spare parts, and service.

· Price protection.

1.2.7
Acceptance Testing

All medical equipment must be inspected and tested before it is placed into service regardless of whether it is purchased, leased, rented, or borrowed by the hospital.  Acceptance testing involves the following:

· An incoming inspection designed to verify that each medical device received is capable of performing its designed function and is electrically safe.

· Baseline measures that may be later used to resolve specified problems.

· Compliance with the equipment assets management process.

The last process is the equipment enters the equipment maintenance program, the warranty period is initiated (if applicable), and the testing criteria are established and documented.

1.3
Clinical Team Requirements

In clinical acquisition process, clinical-decision makers analyse the following criteria concerning proposed technology acquisitions, specifically as they relate to clinical team requirements:

· Ability of staff to assimilate the technology.

· Medical staff satisfaction (short-term and long-term).

· Impact on staffing (numbers, functions).

· Projected utilization.

· Ongoing related supplies required.

· Effect on delivery of care and outcomes (convenience, safety, or standard of care).

· Result of what is written in the clinical practice guidelines.

· Credentialing of staff required.

· Clinical staff initial and ongoing training required.

· Effect on existing technology in the department or other services/departments.

2. RISK MANAGEMENT (RM)

2.1
Introduction

The next task of managing technology throughout its expected life, by focusing on improving performance through monitoring quality assurance indicators and reduction of risk potential.

The health care delivery system is comprised of many complex activities, each of which provides opportunities for interactions among people, agents of technology, and interventions. 

In each intervention, there is a level of real risk coupled with some uncertainty.

Risk measures the probability and severity of loss or injury, while uncertainty refers to lack of knowledge or lack of certainty.

While it is not simple to quantify the cost-to-benefit ratio in a particular intervention, risk management evolves into responsibility that includes prediction of injury, avoidance of exposure to risks, and minimization of liability exposure.

A proactive approach to identify and correct problems requires the use of management tools that reduces risk and increase benefit.

Clinical Engineers can implement equipment-related decision making programs that reduce uncertainty and improve quality through the implementation of a proper evaluation and monitoring system.

2.2
Reducing Risk

Significant progress of controlling risks has been achieved with the implementation of equipment assets management programs.

Recent development of dynamic equipment risk factors and the associated failure analysis technique, proactive techniques to contain risk can be implemented.

Physical risks associated with the application of equipment categorizes risk level at a time equipment enters into hospital’s equipment inventory.

The dynamic risk factor modifies this level and individually identifies each piece of equipment within an equipment category according to the accumulated experiences over the equipment life cycle.  

These experiences include:

· number of failures, 

· number of recalls, 

· hazards, or incidents, and 

· equipment age.

A fundamental management commitment to minimizing the adverse effect of accidental loss to an organization is the founding principal of risk management program.  This is the process of making and carrying out decisions that minimize incidents.

The RM program requires:

· the development of risk criteria, 

· the identification of problems, and

· an action plan to reduce them.

The technology management program participates in the organisation’s effort at the beginning and throughout the equipment life cycle by assessing equipment performance.

The impact of both risk and quality are monitored prior to the purchasing decision, during installation, maintenance and repair, and as indicators for replacement.

Faulty design, poor manufacturing, lack of compatibility with existing technology, and mismatch with user skills and/or needs can be corrected during the equipment selection and acceptance testing.

Incorrect operating procedure, lack of or inadequate maintenance program, or faulty repair work can be corrected by failure evaluation system.

The collection of equipment failure analysis information over several years indicates that equipment risk factor is a dynamic phenomenon.

The equipment-related dynamic risk factor is a modification of a static factor, which is assigned, to each type of equipment at the time first enters equipment assets management program.

The static risk factor is modified continually over the equipment’s life cycle by risk factors that are derived from information collected about experiences with each equipment performance.

Periodically, a summary report of significant equipment-related performance is prepared.

The report’s data is made of elements that show:

1. the ratio of the number of completed-to-scheduled inspections,

2. the number and percent of devices failing to pass prescribed inspection,

3. the number and percent of devices for which  a user’s complaint was registered (however no problems were found),

4. the number and percent of devices that show physical damage,

5. devices that were involved with an unusual event (incident), and

6. the number of devices for which a user’s complaint was justified.

Each element is counted as an event, and thresholds above which unsafe conditions may exist can be determined.

This risk assessment and reduction program should be complemented with professional communication regarding on-going equipment performance between clinical engineering, the user’s department, and various manufacturers. 

This will result in the availability of better and safer products, less complex operation and maintenance instructions, more effective in-service training, and rapid resolution of equipment problems when action is needed.

2.3

Quality Improvement (QI)

A quality improvement process is introduced, in order to provide a patient care environment that is continuously safer and provides higher quality patient outcomes.

It is a philosophy and a methodology focusing on continual assessment of performance and improvement of clinical, operational, administrative and other hospital processes.

One of the QI process methods is based on 10 steps that are the building blocks for an information collection and evaluation system – measurement and assessment of performance – that allows managers to focus, through an organized system, on improving that performance.

The process is proactive and consists of:

1. structure – written manual of policies and procedures defining goals and objectives,

2. process – understanding of the operation and the ability to determine opportunities for improvement, and

3. outcome base – an evaluation of the program impact as determined by function reports.

Problems are considered outliers and, through appropriate selection, can be order ranked.  Selectively, the top 10% to 20% can be addressed at a time.

This is a continuous structured approach that improve the quality of the technology management program by setting minimum acceptable performance goals, by increasing equipment availability for intended clinical use, by ensuring optimal device function, and by eliminating device-related incidents.

The 10 steps are:

Step 1:
Assign Responsibility

Step 2:
Delineate Scope of Service

Step 3:
Identify Important Aspects of Service

Step 4:
Identify Indicators

Step 5:
Establish Thresholds for Evaluation

Step 6:
Collect and Organize Data

Step 7:
Evaluate Service

Step 8:
Take Actions to Solve Problems

Step 9:
Assess Actions and Document Improvement

Step10:
Communicate Relevant Information

2.4

Codes and Standards

Risk analysis, continuous QI, cost/benefit assessment, and compliance with the regulatory mandates are concepts that are vital for manager who is responsible for managing medical technology.

Technology introduces, extends, modifies, or eliminates real as well as perceived risks.  Its management requires a mixture of subjective judgement and objective measurement that at times may get lost in the alphabet of regulatory jargon.

For example, IEC (International Electro-technical Commission), Safe Medical Devices Act (SMDA), American National Standards Institute (ANSI), Medical Information Bus (MIB), Joint Commission on Accreditation of Healthcare Organisation (JCAHO), National Institute for Occupational Safety and Health (NIOSH).  This is sample of the many that mandates and activities that shape the administration of healthcare services.

Improvement in patient care, clinical performance, and desired outcome are critical elements of a system where codes and regulations are continuously evolving. 

It is important for the technology manager to understand fully how standards are developed, how they are used and modified, and most significantly, the effect of these activities on the entire spectrum of health related matters.

PAGE  
1
GSR/KBEB 3234 Perspective in Healthcare Technology/Topic3-Equipment Acquistion


