
Replacing and Adjusting Fluorescent Bulb

Replacing Bulb:
	HBO 100 W/2

[image: image1.jpg]S

de

Back

- Housing

Up/Down

) #1
\ ; fHousing
0
/) \ Remove
Mirror Housing
Front/Back Rotate






	HBO 100

[image: image2.jpg]Side

#1
Housing
Remove

< \J

\
Moves Moves

Bulb Mirrors







HBO 100 W/2:
□ Unplug power supply
□ Remove housing from microscope by unscrewing (c-mount)
□ Take off casing by unscrewing #1 on diagram
** Create space for the bulb to exit by retracting the mirror all the way (clockwise) and by retracting the condenser (#2) 
□ Bulb is held in heat sink with a screw, remove it before removing bulb.
□ DO NOT TOUCH GLASS OF BULB – If it is touched, it must clean with the provided alcohol wipe
□ Replace bulb (small end into heat sink) and tighten the screw firmly
□ Reassemble light housing
□ Reset hours
HBO 100:

□ Unplug power supply
□ Remove housing from microscope by unscrewing (c-mount)
□ Take off casing by unscrewing #1 on diagram

□ Remove bulb/heat sink by pressing down on lever behind mirror and pull bulb out vertically

□ Bulb is held in heat sink with a screw, remove it before removing bulb.
□ DO NOT TOUCH GLASS OF BULB – If it is touched, it must clean with the provided alcohol wipe
□ Replace bulb (small end into heat sink) and tighten the screw firmly
□ Reassemble light housing
□ Reset hours
Adjusting:
□ Remove objective
□ Place a business card or the white label of a cover glass to see fluorescent beam
□ Move housing (HBO 100 W/2) or bulb (HBO 100) to position with maximum brightness

□ Should see two circular dots, one of which moves with the mirrors.

□ Move the mirrors so that the dots overlap.
□ Adjust condenser (#2) to get a uniform diffuse or sharp dot.
□ Replace objective and fine tune the housing (HBO 100 W/2) or bulb position (HBO 100) so that the fluorescent beam is centered.

http://www.geocities.com/simonwaynemoore/protocols.htm


 Page 1 of 1

